WEBSITE COPY REMEDIATION DEBRIEF

SonoHL Website Alignment

Date: June 23, 2026  |  Commit: 486f22c  |  Source: A0-audited whitepaper v0.1 DRAFT

Outcome. The entire website narrative was changed from promotional medical-product copy to controlled investigational-development language. Unsupported diagnostic, performance, configuration, trial-status, remote-readiness, and value claims were removed or recast as hypotheses. The visual structure and existing single-file implementation were preserved.

What Changed

Area

Remediation

Reason

Product positioning

Reframed SonoHL as an “investigational acoustic-sensing platform in development.” Added that the product is not available for commercial sale.

The intended use, configuration, clinical output, and regulatory pathway are not yet approved.

Diagnostic and screening claims

Removed “AI-powered screening,” diagnostic screening, defect detection, and diagnostic/telemonitoring product positioning.

These claims require an approved intended use, clinical evidence, and applicable regulatory authorization.

Performance claims

Removed claims of automatic clearest-signal selection, real-time analysis, noise reduction, clearer recordings, and abnormal-sound detection.

No controlled specifications, acceptance criteria, verification reports, or performance results were supplied.

Device configuration

Removed factual 16-channel, 10-front/6-back, 8+, and 8-sensor statements. Diagrams are now labeled illustrative.

The source materials contained conflicting configurations; the current controlled device version is unresolved.

Epidemiology and trial status

Removed the 2.5 million, 40%, and “CT1 currently enrolling” statements.

The audit did not verify exact primary sources or a current controlled trial record.

Remote use

Recast telemonitoring and home/remote use as future evaluation concepts with explicit dependencies.

Remote readiness requires validated setup, transmission, privacy, cybersecurity, review, and escalation workflows.

Clinical workflow

Changed the workflow to prepare/fit, acquire/transfer, and qualified human review. Added controls still required at each stage.

The original simplified workflow implied operational readiness and established clinical use.

Evidence page

Replaced unsupported development-status claims with a technical, clinical, human-factors, and workflow evidence framework.

Prototype, validation, pilot, and regulatory statuses require verified records and approvals.

Team and governance

Marked titles, affiliations, roles, and governance wording as pending verification.

The audit did not independently confirm current titles, institutional wording, or publication permission.

Partnership outreach

Changed promotional investment/diagnostic language to controlled product and evidence-development collaboration.

External outreach must disclose investigational status and distinguish objectives from demonstrated results.

Representative Before and After

Previous copy

Updated copy

“AI-Powered Cardiopulmonary Screening”

“Investigational Acoustic-Sensing Platform”

“Records heart sounds from 16 channels… automatically selects the clearest signals… analyzes them in real time…”

Captures sounds from multiple predetermined positions and organizes recordings for clinician review; exact configuration and outputs remain under development.

“Heart and Lung Diagnostics. One Vest.”

“Acoustic assessment and remote review concepts,” explicitly identified as development hypotheses.

“Signal processing reduces noise and organizes the clearest recordings.”

Specified quality checks and processing steps are to identify usable and unusable recordings.

“Clinical trial currently enrolling.”

Trial and development status require verified protocols, ethics status, sites, enrollment records, and approval.

Why This Is a Material Improvement

•	
      Reduced regulatory exposure: the public narrative no longer presents investigational objectives as established safety, effectiveness, diagnostic, or commercial facts.

•	
      Improved internal consistency: conflicting sensor counts and development stages were replaced with controlled-input requirements.

•	
      Clearer product boundaries: the site now consistently states that human review is required and that SonoHL does not independently diagnose disease or replace medical judgment.

•	
      Better evidence discipline: technical validation, clinical validity, human factors, privacy, cybersecurity, and workflow evidence are separated rather than blended into promotional promises.

•	
      More defensible partnership language: collaborators are invited to support evidence generation without implying market readiness or demonstrated clinical benefit.

Validation Performed

•	
      Reviewed every visible website text block, including headings, labels, cards, hidden fragments, SVG labels, navigation, calls to action, and image alternative text.

•	
      Scanned visible content for quarantined phrases covering AI, diagnosis, screening, detection, real-time processing, noise reduction, sensor counts, epidemiological statistics, and trial enrollment; no blocked visible claims remained.

•	
      Confirmed all nine site sections still parse and remain present.

•	
      Ran Git whitespace/diff validation and checked navigation fit after shortening revised labels.

•	
      Committed and pushed only sonohl.html; unrelated local documentation files were not included.

Remaining Actions

•	
      Regulatory and legal review of the exact investigational-device wording for each publication channel.

•	
      Confirm current team titles, clinical-governance roles, affiliations, and permission to publish.

•	
      Freeze the controlled hardware/software configuration and replace illustrative labels only after approval.

•	
      Supply controlled clinical-program records before publishing status, trial, performance, or regulatory-pathway statements.

•	
      Run the planned follow-up audit against the live rendered site and all other collateral.

Scope note: This remediation changed website copy and accessibility labels only. It did not refactor the HTML, CSS, JavaScript, embedded images, hosting, analytics, or underlying product implementation.